## Sheet1
Exhibit: X
Lab Type: APF Clean Room
Project: | [8 empty] | Room Name:
WR Number: | [8 empty] | Room Number:
Date:
Room Data | [8 empty] | Other Special Requirements
Size/Dimensions | (1) module, 11’ width | [7 empty] | Anteroom required for gowning with bench, PPE storage, waste bins
PEC | <797> Isolator
Ceiling height | 2,896 mm (9’-6”) minimum
Door size | 1,200 mm (4’-0”) | [7 empty] | Active leaf 900 mm (3’-0”); inactive leaf 300 mm (1’-0”)
Door type | Cleanroom unequal pair, full lite (active)/Solid (inactive) | [7 empty] | Card key access control
Access Control | Red-Light/Green-Light Indicator
Windows | Cleanroom Window
Normal occupancy | 4
Special requirements | Confirm with URS | [7 empty] | All surfaces cleanable; sticky mats at entrance; all penetrations sealed
2. Finishes | [8 empty] | Other Special Requirements
Floor | Seamless Sheet Vinyl, Welded
Base | 6” integral with floor | [7 empty] | Reinforced, factory formed coved corners
Wall type | Cleanroom modular panelized wall system | [7 empty] | Bubble-Tight stainless steel access panel, where required
Wall Protection | Stainless steel (304L or 316L) | [7 empty] | Full-height corner guards, and fully adhered scuff plates, gapped 1/2"
Ceiling type | Cleanroom modular panelized ceiling system | [7 empty] | Bubble-Tight stainless steel access panel, where required
3. Furnishings and Fittings | [8 empty] | Other Special Requirements
Casework | Stainless steel (304L or 316L) | [7 empty] | Minimize casework to promote cleanliness; use tables and carts where possible
Bench top | Stainless steel (304L or 316L)
Sink(s) | No
Piped utilities | No
Flammable storage cabinet | No
Vented corrosive storage cabinet | No
Other | Minimize horizontal surfaces, cracks & seams
4. Equipment – See Equipment List for Additional Items | [8 empty] | Other Special Requirements
Biological Safety Cabinet (BSC) | (1) 6’, class II, type A2 | [7 empty] | Vacuum in BSC
Laminar Flow Hood (LFH) | No
Compounding Aseptic Isolator (CAI) | No
Containment Isolater (CACI) | No
Compounding Safety Enclosure (CSE) | No
Restricted Access Barrier System (RABS) | No
Controlled Rate Freezer (CRF) | No
Laboratory Freezer (-20/-30/-80) | No
Incubator | No | [7 empty] | Water Jacketed, CO2, N2, and air
Microscope | Yes | [7 empty] | Phase Contrast, Photo Port
Other
5. HVAC Requirements | Min Regulatory Value |  |  | Min. Alarm Setpoint |  |  | Setpoint |  |  | Max. Alarm Setpoint |  |  | Max. Regulatory Value | Alarm Delay
Temperature: | 25°C (59° F) |  |  | 16°C (60.8°F) |  |  | 20°C (68°F) |  |  | 21°C (69.8°F) |  |  | 22.2°C (72°F) | 2-Min
Relative Humidity: | N/A | N/A |  | N/A |  |  | 40% RH |  |  | 42% RH |  |  | 50% RH | 2-Min
Differntial Pressure: | 0.020 IWC | N/A |  | 0.025 IWC |  |  | 0.033 IWC |  |  | 0.035 IWC |  |  | N/A | 2-Min
6. HVAC System Description | Notes
ISO Class | ISO-7
Airflow Type | Unidirectional
Supply Air | 472 LPS (1,000 CFM) Ceiling HEPA diffusers
Exhaust Air | 377.6 LPS (800 CFM) low sidewall grille, 47.2 LPS (100 CFM) door undercut
Recirculation Air | 472 LPS (1,000 CFM) Ceiling HEPA diffusers
Relative Pressure | 12.5 Pa (0.05 " WC) positive to anteroom
BAS Sensors | TEMP, RH, dP, and Door Position Red/Green Light
EMS Sensors | TEMP, RH, and dP
6. Piping | CHW | CW | PW | HW | OTHER | CO2 | AIR | VAC | GAS | WASTE | STEAM | FD
Utility | No | No | No | No | No Condensate | Yes | No | Yes | No | No | No | No
Other | Piping services as determined by program
7. Electrical |  | Other Special Requirements
Power receptacles | Yes | [7 empty] | General purpose NEMA 5-20R receptacles, cast boxes and conduits sealed
Lighting | Recessed Cleanroom LED Luminaire | [7 empty] | Lensed, sealed, gasketed, wet location listed, with overlapping door
Telephone/Communication | Yes
Data/Computer | Yes
Emergency power | Yes | [7 empty] | Emergency power for equipment per equipment list
Standby Power | No
Task lighting | Yes